[image: image1.jpg].

svaraan eurex co.i. =N

10000 Matton sty 2o, anan i snarasan cusnason- raoss . o =M1 T EK
e ko 1oL 8675209 300 T80 TES-SRBS 4008 e 35 120 ot

o, Lazer

| Certificate of Conformity

NO.:KA0811012

The following product has been tested by us with the listed standards and found in conformity with the

| requirements.
I Applicant : AURORA ELECTRONIC(CHENGDU)CO., LTD
Address + No. 50, Tongsheng Road, Qingyang District, Chengdu,
China
Manufacturer : AURORA ELECTRONIC(CHENGDU)CO., LTD
Address + No. 50, Tongsheng Road, Qingyang District, Chengdu,
China
EUT + Carbon Monoxide Detector
MN : CO-011,CO-012, CO-021, CO-022, CO-031, CO-032,
€0-200D, CO-C200E
Technical Data : 9V-12VDC/1 10-230VAC, S0/60Hz
Test Standards : EN50291: 2001

(Manager)
November 17, 2008

Shanzne®

The test report was carried out from the submitted type samples of a product in conformity with the
specification of the respective standards.





sira Sira Environmental Ltd
12 Acorn Industrial Park Crayford Road Crayford,Dartford Kent, DA1 4AL  01322 520509  01322 520501 : richard.king@siraenvironmental.com
	Test
	 Test result
	 BS EN 50291 specification

	 
	 
	 

	 
	 
	 

	 ALARM CONDITIONS
	 
	 

	 
	Sample Number
	The alarm shall not operate
during or after exposure to
33 ppm CO for 120
minutes.  The alarm shall
operate after between 60
and 90 minutes on
exposure to 55 ppm CO,
between 10 and 40 minutes on exposure to 110 ppm 
CO, and within the first
3 minutes of exposure to
330 ppm CO.
Recovery from the alarm 
state shall take place,
after manual resetting if
necessary, within 6
minutes when exposed to
clean air.

	Time to alarm, after warm-up:
	4
	5
	6
	

	 
	 
	

	            33 ppm CO
	No alarm
	No alarm
	No alarm
	

	            
            55 ppm CO
	 
68 min
	 
68 min
	 
68 min
	

	 
	 
	 
	 
	

	            110 ppm CO
	26 min
	29 min
	38 min
	

	 
	 
	 
	 
	

	            330 ppm CO*
	29 s
	2 min 00s
	2 min 32s
	

	 
	 
	 
	 
	

	Maximum time for alarm to cancel on exposure to clean air
	5 min 10s
	6 min†
	6 min
	

	 
	 
	

	 
	 
	 
	 
	

	 
	 
	

	 
	 
	 
	 
	

	 
	 
	 
	 
	

	 
	 
	 
	 
	

	 
	 
	 
	 
	

	
	 
	


NOTE:-
 
*          Includes an allowance for the flush time of the test box.
†          Manual reset of unit.
 
The units appear to comply with Clause 5.3.4 of EN 50291. The test report, samples and invoice will be sent to you on Monday.
 
Best regards,
 
Richard
 
Richard King
Instrument Test Engineer 
 
 

